BRA-Logger: AI-Powered Breast Anomaly Risk Scoring System

Complete Implementation Plan

Project Goal: Develop an IoT-based wearable smart bra system that uses embedded temperature, pressure, and bioimpedance sensors to capture breast tissue data, process it through an AI/ML pipeline, and generate anomaly risk scores, assisting in the early identification of abnormal breast tissue patterns.

Table of Contents

- Problem Statement & Clinical Rationale

- Data Strategy. Datasets & Collection

- Hardware Requirements. Complete Bill of Materials

- Tech Stack

- Project Structure. Folder Architecture

- Software Implementation. Phased Breakdown

- ML/AI Pipeline Design

- API Design

- Frontend Dashboard Design

- Deployment & DevOps

- Testing Strategy

- Ethical, Regulatory & Legal Considerations

1. Problem Statement & Clinical Rationale

The system measures three biological signals (biomarkers) from breast tissue:

- Temperature

- Pressure / stiffness

- Electrical impedance

Abnormal breast tissue behaves differently from normal tissue, and those differences can be measured with sensors.

Why Temperature & Pressure?

Cancerous and pre-cancerous breast tissue exhibits measurable physiological differences

The table shows scientific evidence that abnormal breast tissue behaves differently from healthy tissue.

Each row = one biological signal (biomarker) that the system measures.

Each column = explanation of what that signal looks like in healthy vs abnormal tissue.

The table structure:

| Column | Meaning |
| --- | --- |
| Biomarker | The biological signal being measured |
| Normal Tissue | What that signal looks like in healthy breast tissue |
| Abnormal/Cancerous Tissue | What happens to that signal when abnormal tissue exists |
| Clinical Basis | Scientific/medical evidence explaining why this happens |

| Biomarker | Normal Tissue | Abnormal/Cancerous Tissue | Clinical Basis |
| --- | --- | --- | --- |
| Surface Temperature (skin temperature over the breast) | Symmetric across both breasts (±0.3°C). Both breasts should have almost the same temperature. Difference between left and right should be very small. ±0.3°C difference | Localized elevated temperature (0.5–2°C higher) due to increased angiogenesis and metabolic activity. If abnormal tissue exists, temperature increases in that area. For example: Left breast: 36.4°C Right breast: 37.3°C Difference = 0.9°C That is suspicious. Why? Because tumors need more blood and energy. So the body sends more blood vessels to that area. More blood = more heat. This process is called: angiogenesis Meaning: formation of new blood vessels | Breast thermography literature (Gautherie, 1980; Kontos et al., 2011). These are thermography studies showing cancer areas are hotter. |
| Tissue Stiffness (Pressure). The system measures how hard or soft tissue feels when pressure is applied. This mimics what doctors do during a clinical breast exam. Doctors press the breast to check for lumps. Our device uses pressure sensors instead of fingers. | Normal breast tissue is Soft, uniform, elastic When pressure is applied, the tissue compresses evenly Meaning sensors detect similar pressure resistance everywhere. | Hard, non-uniform lumps with higher stiffness. Tumors are usually hard, dense, stiff They resist pressure. So sensors detect higher pressure resistance | Clinical breast exam (CBE) principles, suggest that tumors are typically 5–10x stiffer than surrounding tissue |
| Bioelectrical Impedance. Your body conducts electricity slightly. Different tissues conduct electricity differently. The device sends a very small safe electrical signal. Then it measures how easily the signal travels through tissue. | Normal tissue has a Higher impedance (~600–1000 Ω at 50 kHz) Ω = ohms (unit of electrical resistance) Higher resistance means, electricity flows less easily. | Lower impedance (~200–500 Ω) due to increased vascularity and cellular disruption. Cancer tissue has: lower impedance Values: 200–500 Ω Why? Cancer tissue has: - more fluid - more blood vessels - damaged cell membranes All of these increase conductivity. So electricity flows easier → impedance drops. | Bioimpedance spectroscopy (BIS) research (Zou & Guo, 2003) |

What We're Detecting

We do not diagnose cancer. We compute an anomaly risk score (0–100) based on statistical deviations from a baseline "healthy" pattern across multiple modalities. This score is a screening aid, not a diagnostic tool.

Why Multiple Sensors Are Used

Our system combines three signals.

Because one signal alone is unreliable.

Example:

| Signal | Possible False Cause |
| --- | --- |
| Temperature | exercise |
| Pressure | natural dense tissue |
| Impedance | hydration |

But if all three change together, it becomes more suspicious.

Example anomaly pattern:

hotspot + stiffness + impedance drop

That pattern is much more meaningful.

2. Data Strategy. Datasets & Collection

2.1 The Dataset Problem

We do not have a pre-built dataset of temperature + pressure + bioimpedance sensor readings from wearable devices. This is novel research, and such datasets don't exist publicly.

There is no public dataset containing:

- breast temperature sensors

- pressure sensors

- bioimpedance sensors

- collected from a wearable device

So, this project proposes three strategies to solve the data problem.

2.2 Strategy A : Synthetic + Simulation Data (Fake data (simulation))

This is the starting point of the project.

We generate fake but realistic sensor data using software.

Why?

Because before the hardware even exists, we’ll still need data to:

- build the ML pipeline

- test the algorithms

- develop the backend

- train early models

So, we have to simulate sensor readings.

The script receiver_logger.py will collect simulated signals.

The table below explains how much synthetic data should be generated.

Column meaning:

| Column | Meaning |
| --- | --- |
| Parameter | What part of the dataset we are defining |
| Value | Recommended amount or setting |
| Rationale | Why that amount was chosen |

| Parameter | Value | Rationale |
| --- | --- | --- |
| Healthy baseline samples | 500–1,000 sessions × 5 min each. 500–1,000 sessions and Each session lasts 5 minutes. What is a session? A session is one data collection event. Example: User wears device for 5 minutes Sensors record signals That is one session Why 500–1000 sessions? Because normal body signals naturally vary. Things that affect readings: - time of day - exercise - menstrual cycle - room temperature - posture So, we need many examples of normal behavior. | Captures normal variation across time-of-day, activity level, menstrual cycle phase |
| Simulated anomaly samples | 100–200 sessions These are fake abnormal cases. Because real cancer data does not exist yet. So the simulator artificially creates anomalies. Example anomaly types: - Temperature Hotspot - Pressure lump - Impedance drop - Left-right asymmetry This helps test whether the AI can detect abnormal patterns. | Injected hotspots, asymmetries, stiffness zones |
| Sampling rate | 2 Hz (existing) Hz means, samples per second So, 2 Hz = 2 measurements per second Example: If session is 5 minutes 5 minutes = 300 seconds Then: 300 × 2 = 600 sensor readings Why is this enough? Because breast tissue signals change slowly. This is not like heart rate or ECG. Temperature and tissue stiffness change gradually. So high-frequency sampling isn't needed. | 2 samples/second is sufficient for slow-changing physiological signals |
| Channels | 4 temperature + 8 pressure + 4 impedance = 16 channels Meaning: Each timestamp records 16 measurements. Example: - Time, t temp1 temp2 temp3 temp4 - press1 press2 press3 press4 press5 press6 press7 press8 - imp1 imp2 imp3 imp4 Each is one channel. Channels represent sensor positions on the breast. | Extended to 16 to include bioimpedance |
| Minimum dataset size for MVP | ~50,000 windows This comes from sliding window processing. Sliding Window Explanation: Raw data is continuous. But ML models work on chunks of time. For Example: 30 second window And windows move every: 5 seconds Example timeline: 0–30 sec 5–35 sec 10–40 sec 15–45 sec This creates many overlapping windows. So even a few sessions produce thousands of training samples. | Enough for Isolation Forest or autoencoder training |

How to simulate anomalies:

- Temperature hotspot: Add +0.5 to +2.0°C to 1–2 specific temp sensors for a sustained period.

For Example: Normal temperature = 36.5°C

Simulated anomaly: +0.5°C to +2°C

For Example: from 36.5 → 38°C Only on specific sensors.

These mimic localized tumor heat.

- Pressure lump: Increase 1–2 pressure channels by 30–60% (simulating a hard mass)

Simulated by increasing pressure sensor values: 30–60%

Why? Tumors are harder than normal tissue. So, sensors detect stronger resistance.

- Impedance drop: Reduce impedance on 1–2 channels by 40–60%

Impedance reduced by: 40–60%

This simulates cancer tissue conductivity.

- Asymmetry: Create a persistent L-R difference exceeding 1.0°C or 25% pressure differential.

Difference between left and right side.

For Example:

Left temp = 36.5°C

Right temp = 37.8°C

Difference: 1.3°C

Which exceeds the 1°C asymmetry threshold.

2.3 Strategy B. Public Datasets for Transfer Learning

These cannot be used directly (different modality), but are valuable for validating feature engineering and model architectures

This means Using existing medical datasets, even though they are not identical

Why?

They contain related information.

Mostly thermal breast imaging.

| Dataset | Source | Content | Use Case |
| --- | --- | --- | --- |
| DMR-IR (Database for Mastology Research) | Visual Lab UFPE | 293 patients Infrared breast images Mammograms Clinical records | Thermal pattern analysis Even though it's images, not sensors. You can still learn: temperature distribution patterns |
| Mendeley Breast Thermography | Mendeley Data | 119 women FLIR thermal camera images labels: benign / malignant | Training models to detect thermal anomalies. The model can learn: What abnormal temperature distribution looks like Even though sensors are different. |
| Kaggle Thermal Breast | Kaggle | Thermal images Categories: normal sick unknown | Use case: Rapid deep learning experiments Example: Train a CNN to detect heat signatures. |
| UCI IR Temperature | UCI ML Repository | Temperature measurements Oral temperature Demographics | Understanding normal body temperature variation. This helps calibrate: What temperature ranges are normal |

2.4 Strategy C — Real-World Data Collection Protocol (Post-Hardware)

Once the hardware exists, we shall collect real sensor data. This is the most valuable dataset. But also the hardest.

Once the wearable is built:

- Healthy Baseline Collection (Phase 1 priority):

  - We’ll have to Recruit 20–50 healthy volunteers

  - Each wears the device for 5 min in 3 conditions: rest, standing, post-light-walk

  - Collect 3 sessions per participant across different times of day

  - Label: healthy, with metadata: age, BMI, menstrual phase, ambient temperature

20–50 volunteers. They wear the device for 5 minutes in 3 conditions; Resting, Standing and After light walking, Because body signals change with activity.

There should be Multiple Sessions, meaning, each participant records 3 sessions at different times of day, because Temperature varies during the day.

Along with sensor data, we’ll collect

| Data | Reason |
| --- | --- |
| Age | physiology differences |
| BMI | tissue composition |
| Menstrual phase | hormonal changes affect temperature |
| Ambient temperature | affects skin temp |

- Known-Anomaly Collection (Phase 2, if clinical access available):

  - Partner with a clinic or hospital

  - Collect readings from patients with known benign lumps or confirmed malignancies

  - Label: benign, malignant, post-surgery

- Minimum Viable Dataset:

Columns

| Column | Meaning |
| --- | --- |
| Category | Type of data |
| Sessions | number of recordings |
| Per-session duration | recording length |
| Total windows | number of ML training samples |

| Category | Sessions | Per-Session Duration | Total Windows (~30s, 5s step) |
| --- | --- | --- | --- |
| Healthy | 150 sessions | 5 minutes | ~7,500 windows |
| Simulated Anomaly | 100 | 5 mins | ~5,000 |
| Real Anomaly (stretch) | 20+ | 5 mins | ~1,000 |
| Total MVP | 270+ |  | ~13,500+ |

Recommended Approach: We’ll Start with Strategy A (simulation). Use Strategy B for model architecture validation. Move to Strategy C once hardware is functional and tested.

3. Hardware Requirements — Complete Bill of Materials

We need to Look at the hardware as five subsystems:

1️⃣ Brain (microcontroller)
2️⃣ Sensors (temperature, pressure, impedance)
3️⃣ Power system (battery + charging)
4️⃣ Supporting sensors (environment + motion)
5️⃣ Assembly materials (wires, boards, enclosure)

3.1 Core Controller. The brain of the device

This is the main computer inside the wearable.

| Component | Specification | Quantity | Purpose | Est. Cost (USD) |
| --- | --- | --- | --- | --- |
| ESP32-WROOM-32D DevKit | Dual-core 240 MHz, Wi-Fi + BLE, 4 MB flash | 1 | It controls everything: - reads sensor data - processes signals - sends data to the computer/server - controls Wi-Fi or Bluetooth | $5 - $8 |
| USB-C cable | For programming & power | 1 | Flashing firmware, serial debug | $3 |

Sensors → ESP32 → Wi-Fi → Laptop → AI model

Key capabilities:

| Feature | Meaning |
| --- | --- |
| Dual-core 240 MHz CPU | small but powerful processor |
| Wi-Fi + Bluetooth | wireless data transmission |
| 4 MB flash | stores firmware |
| ADC inputs | reads analog sensors |

Why ESP32 was chosen

Because it is:

- cheap

- powerful

- built-in Wi-Fi

- widely supported

Cost:

$5 – $8

You also need a USB-C cable to program it.

3.2 Temperature Sensors

| Component | Specification | Quantity | Purpose | Est. Cost (USD) |
| --- | --- | --- | --- | --- |
| DS18B20 digital temperature sensor | Digital, ±0.5°C accuracy, 1-Wire protocol Sensor accuracy is ±0.5°C Which is enough because anomalies are usually: +0.5°C to +2°C The sensor uses something called 1-Wire communication. This means: All temperature sensors share ONE wire. So wiring is simple. We’ll also need a 4.7 kΩ resistor. Purpose: stabilizes the data line | 4 | These measure skin temperature on the breast. Placement: 2 per breast Positions: Upper outer quadrant Lower inner quadrant | $4 - $6 (a pack of 5) |
| 4.7 kΩ resistor | Pull-up resistor for 1-Wire data line | 1 | Required by DS18B20 1-Wire protocol | $0.10 |

Why this sensor was chosen

Advantages:

- digital (no noise)

- cheap

- reliable

- multiple sensors on one pin

Upgrade option: ADT7420 temperature sensor

More accurate: ±0.25°C But more expensive.

3.3 Pressure/Force Sensors

| Component | Specification | Quantity | Purpose | Est. Cost (USD) |
| --- | --- | --- | --- | --- |
| FSR 402 Force Sensing Resistor | 0–10 kg, 12.7mm active area | 8 | These detect tissue stiffness. This mimics what a doctor does during a breast exam. When you press the breast: - soft tissue compresses - tumors resist pressure The sensor measures that resistance. Placement: 4 sensors per breast Positions like a clock: 12 o’clock 3 o’clock 6 o’clock 9 o’clock This creates full spatial coverage. How they work FSR sensors change electrical resistance when pressure is applied. Example: no pressure → high resistance more pressure → lower resistance The ESP32 reads this as voltage change. | $16–24 (pack of 8) |
| CD74HC4067 analog multiplexer | Voltage divider for each FSR | 1 | Problem: The ESP32 has few analog inputs. But we have many pressure sensors. Solution: The multiplexer acts like a sensor switchboard. Many sensors → Multiplexer → ESP32 The ESP32 reads sensors one at a time very quickly. like: sensor1 sensor2 sensor3 sensor4 sensor5 sensor6 sensor7 sensor8 Instead of needing 8 wires to ESP32, the multiplexer lets them share. | $3 |

3.4 Bioimpedance Sensors

| Component | Specification | Quantity | Purpose | Est. Cost (USD) |
| --- | --- | --- | --- | --- |
| AD5933 impedance converter | Breakout board (I2C), 1 kHz–100 kHz sweep | 1 | It measures electrical impedance of tissue. The device sends a tiny electrical signal through the tissue. Then measures: how easily electricity flows Cancer tissue typically has: lower impedance because of: - higher blood flow - damaged cell membranes - increased fluid | $15 - $20 |
| Ag/AgCl gel electrodes (disposable) | Pre-gelled, snap-on, medical grade | 8 (4 pairs) | These stick to the skin. They are the same type used in: - ECG tests - hospital monitors They attach with snap connectors. Measurement method The system uses: 4-electrode measurement Why? Because it improves accuracy. | $5 (pack of 50) |
| Electrode snap connectors | Standard 3.5mm or 4mm snap | 4 | Connect electrodes to AD5933 board | $3 |

3.5 Additional Sensors

| Component | Specification | Quantity | Purpose | Est. Cost (USD) |
| --- | --- | --- | --- | --- |
| DHT22 / BME280 | Ambient temp + humidity sensor These measure: room temperature humidity Why? Because skin temperature depends on environment. Example: cold room → skin temp drops So, the system adjusts readings. | 1 | It measures electrical impedance of tissue. The device sends a tiny electrical signal through the tissue. Then measures: how easily electricity flows Cancer tissue typically has: lower impedance because of: - higher blood flow - damaged cell membranes - increased fluid | $3–$5 |
| MPU6050 / MPU9250 IMU) | Accelerometer + gyroscope (I2C) | 1 | This measures: movement orientation Why important? Because movement causes sensor noise. Example: walking bending breathing motion The AI can filter those artifacts. | $3–5 |

3.6 Power

| Component | Specification | Quantity | Purpose | Est. Cost (USD) |
| --- | --- | --- | --- | --- |
| 3.7V LiPo battery | Capacity: 1000 – 2000 mAh JST connector | Portable power for wearable operation | It measures electrical impedance of tissue. The device sends a tiny electrical signal through the tissue. Then measures: how easily electricity flows Cancer tissue typically has: lower impedance because of: - higher blood flow - damaged cell membranes - increased fluid | $5–8 |
| TP4056 micro-USB charging module | Li-ion charger with protection circuit | 1 | Safe LiPo charging | $1 - $2 |
| Slide switch (SPST) | Mini toggle or slide | 1 | A simple slide switch to turn the device on/off. | $0.50 |

3.7 Wiring & Assembly

| Component | Quantity | Purpose | Est. Cost |
| --- | --- | --- | --- |
| Breadboard (for prototyping) | 1 | Initial circuit prototyping | $3 |
| Jumper wires (M-M, M-F, F-F assorted) | 40-pack | Sensor connections during prototyping | $3 |
| Perfboard / Stripboard | 1 | Soldered prototype (post-breadboard) | $2 |
| Header pins (male + female) | 2 strips | Modular connections on perfboard | $1 |
| Heat shrink tubing assortment | 1 pack | Insulate solder joints | $2 |
| Soldering iron + solder wire | 1 set | Assembly (if not already owned) | $15-$25 |
| Hot glue gun | 1 | Sensor mounting on bra fabric | $5 |
| Flexible ribbon cable / thin gauge wire | 1m | Route sensor wires inside bra fabric channels | $2 |
| 3D-printed enclosure (optional) | 1 | House ESP32 + battery + MUX neatly | $5-$10 (filament) |
| Sports bra (for embedding) | 1 | Physical housing for sensors | $10-$15 |

These are the basic electronics tools:

| Item | Purpose |
| --- | --- |
| Breadboard | testing circuits |
| Jumper wires | connections |
| Perfboard | permanent soldered board |
| Header pins | detachable components |
| Heat shrink | wire insulation |
| Soldering iron | assemble electronics |
| Hot glue gun | mount sensors in bra |
| Flexible ribbon cable | run wires through fabric |

Physical Structure

The device is embedded inside a sports bra

Sensors are mounted inside the fabric.

The electronics (ESP32 + battery) go inside a small enclosure.

Optional: 3D printed case

Estimated Cost

| Category | Cost |
| --- | --- |
| Core electronics | $30 – $50 |
| Sensors | $40 – $55 |
| Power system | $7 – $11 |
| Assembly materials | $25 – $45 |

Total prototype cost: $100 – $160

Hardware Communication Architecture

The system flow looks like this:

Sensors
   ↓
ESP32
   ↓
Wi-Fi
   ↓
Laptop / Raspberry Pi
   ↓
Data Logger (Python)
   ↓
CSV files
   ↓
Machine Learning pipeline
   ↓
Backend API
   ↓
Frontend dashboard

3.8 Hardware-to-Software Communication Architecture

So hardware only does the data acquisition and All AI happens on the server.

The Real Hardware Design Idea

The device collects three types of physiological signals simultaneously:

Thermal signals
Mechanical signals
Electrical signals

This is called multimodal sensing.

Which makes anomaly detection much more reliable.

Hardware summary

Brain:
ESP32 microcontroller

Sensors:
4 temperature
8 pressure
4 impedance

Supporting sensors:
environment + motion

Power:
LiPo battery + charging module

Assembly:
wires, board, bra enclosure

Total:

16 sensor channels

feeding the AI system.

FULL SYSTEM PIPELINE

- From Body To Sensors

The system is not detecting cancer but detecting patterns in signals; temperature, pressure and impedence

| Signal | What it captures | Why it matters |
| --- | --- | --- |
| Temperature | Heat variation | Tumors = higher metabolic activity |
| Pressure | Tissue stiffness | Tumors = harder tissue |
| Impedance | Electrical resistance | Tumors = different conductivity |

So, what we have is Time-series sensor data, for example:

{

"timestamp": 1712345678,

"temp_1": 36.8,

"temp_2": 37.4,

"pressure_1": 0.21,

"pressure_2": 0.67,

"impedance_1": 480,

"impedance_2": 510

}

This is still garbage to a model unless processed.

- Firmware:

This is code running on the microcontroller(ESP32) that turns physical signals into structured data.

Basically writing:

Python (backend dev)

↓

C/C++ (firmware dev)

The microcontroller (ESP32) is supposed to be doing 4 jobs:

- Reading sensors

Each sensor behaves differently

- Temperature (Digital)

[C++]

float temp = ds18b20.getTempC();

- Pressure (Analog)

[C++]

int raw = analogRead(pin);

float pressure = raw * (3.3 / 4095.0);

- Impedance (via I2C)

[C++]

float impedance = ad5933.readImpedance();

ii. Cleaning signals

Basic Filtering:

[C++]

// Moving average filter

filtered = (prev + current) / 2;

iii. Packaging data

This is the structure that the backend expects:

[C++]

String payload = "{";

payload += "\"temp1\":" + String(temp1) + ",";

payload += "\"pressure1\":" + String(p1);

payload += "}";

iv. Send data (WiFi)

[C++]

HTTPClient http;

http.begin("http://your-api/sensor");

http.POST(payload);

- Data Ingestion (FastAPI)

Example Endpoint:

[python]

@app.post("/sensor")

def receive_data(data: SensorData):

save_to_csv(data)

return {"status": "ok"}

IoT → API → Storage

- Data Storage

We’ll use PostgreSQL  for storing time-series data

- Feature Engineering

The raw data is converted into useful features

Temperature Features:

- Mean temperature

- Left vs right asymmetry

- Max temperature difference

Pressure Features:

- Average stiffness

- Variance

- Local peaks (possible lumps)

Impedance Features:

- Mean impedance

- Spatial variation

Example Feature Vector

[Python]

[

temp_diff_left_right,

max_temp,

pressure_variance,

impedance_mean,

asymmetry_score

]

Now, the model has something meaningful

- Machine Learning Model

Now we train:

Input: features

Output: risk score (0 → 1)

Example prediction:

[python]

risk = model.predict_proba(features)[1]

Output:

0.82 → HIGH RISK

0.21 → LOW RISK

- Risk Scoring Logic

The model output is translated as required:

| Score | Meaning |
| --- | --- |
| 0.0 – 0.3 | Low risk |
| 0.3 – 0.6 | Moderate |
| 0.6 – 1.0 | High risk |

This is what users actually see

- Frontend

We display:

- heatmap of breast

- pressure map

- trend over time

- risk score

Example:

Left breast: 0.72 → High risk
Right breast: 0.18 → Low risk

TECH STACK

- Firmware (Embedded)

| Technology | Purpose |
| --- | --- |
| PlatformIO + Arduino framework | Build system and HAL for ESP32 |
| C/C++ | Firmware language |
| OneWire + DallasTemperature libraries | DS18B20 driver |
| Wire (I2C) | AD5933, BME280, MPU6050 communication |
| WiFi + WebServer (ESP32) | HTTP API to stream sensor data |
| ArduinoJson | Structured JSON responses |

- Data Collection and Processing

| Technology | Purpose |
| --- | --- |
| Python 3.10+ | Core language for data scripts, ML, backend |
| Pandas | Data loading, cleaning, windowing |
| NumPy | Numerical computation, feature engineering |
| SciPy | Signal processing (filtering, FFT) |

- Machine Learning

| Technology | Purpose |
| --- | --- |
| Scikit-learn | Isolation Forest, Random Forest, StandardScaler, evaluation metrics |
| XGBoost | Gradient-boosted classification/regression for risk scoring |
| PyTorch (optional Phase 3) | Autoencoder for unsupervised anomaly detection |
| Joblib | Model serialization/deserialization |
| Optuna (optional) | Hyperparameter optimization |
| SHAP | Model explainability — which sensors/features drive the risk score |

- Backend

| Technology | Purpose |
| --- | --- |
| FastAPI | REST API framework (async, high-performance, auto-docs) |
| Uvicorn | ASGI server |
| SQLAlchemy | ORM for database models |
| Alembic | Database migrations |
| PostgreSQL | Production database (structured patient/session data) |
| SQLite | Development/testing database |
| Pydantic | Request/response validation |
| python-jose + passlib | JWT authentication |
| WebSocket (FastAPI) | Real-time streaming of live sensor data |

- Frontend

| Technology | Purpose |
| --- | --- |
| React 18+ (Vite) | SPA dashboard framework |
| TypeScript | Type-safe frontend development |
| Recharts / Chart.js | Time-series charts for sensor data |
| React-Heatmap-Grid or D3.js | Breast heatmap visualization |
| Axios | HTTP client for API calls |
| TanStack Query (React Query) | Server state management and caching |
| React Router | Client-side routing |
| Zustand or Context API | Client state management |
| Tailwind CSS or Material UI | UI styling (user preference) |

- DevOps & Tools

| Technology | Purpose |
| --- | --- |
| Git / GitHub | Version control |
| Docker + Docker Compose | Containerized development & deployment |
| pytest | Backend testing |
| Vitest / Jest | Frontend testing |
| pre-commit | Linting hooks (ruff, black, eslint) |
| GitHub Actions | CI/CD pipeline |